## Slide 1: Compact OCT scanner for a bimodal endoscopic probe
Sergio Vilches
Notes: [Keep whole laser beam]

## Slide 2: Motivation Optical biopsy
Arteriovenous Malformation Fujii T. Kuwano H.
Transversal OCT image of the macula J. A. Flatter and J. Carroll,
http://colorectalsurgeonssydney.com.au/?page_id=425
joimax GmbH
Microscopy
Histology
OCT
Biopsy
Conventional biopsy (ex situ, ex vivo)
Optical biopsy (in situ, in vivo)
Diagnostic
Diagnostic
Notes: Analyze tissue inside body: Endoscope. Camera
Conventional biopsy
  Forceps: extract tissue
  Histology:
Alternative developed in recent years:
  Using advanced optical techniques to diagnose
  Microscopy + OCT
Here IMTEK designing such probe

## Slide 3: Fiber scanner Bimodal implementation: Size
Notes: Get an intuition of the size

## Slide 4: Fiber scanner Bimodal implementation
Notes: Proposed implementation of bimodal probe
2 layer design
  Independent beampaths, combined by dichroic beamsplitter
  Ensure registration of both fields
  Unfold the path
This thesis focuses on the OCT beampath design
-> Brief introduction to OCT

## Slide 5: Swept-Source Optical Coherence Tomography SS-OCT Theory
sample
reference
mirror
Single sample reflector
Multiple reflectors
Notes: SS-OCT = Technique to measure depth information of a sample by analyzing its backscattered light.
Michelson interferometer with a laser which can sweep its wavelength.
Intuitive explanation: Imagine reflector in sample arm
  If placed at same distance as reference arm: constructive interference at all wavelengths
  If placed at z1, some wavelengths constructive, other destructive
  Spectrogram modulated by distance
Cross correlations / Autocorrelations
[Distribute]

## Slide 6: Swept-Source OCT: Fiber scanner requirements
Optical requirements:
Long depth of field: > 2 mm
Scanning mechanism
High field of view for 2 mm lens
Telecentric
Mechanical requirements:
Diameter < 3.3 mm
Scanning mechanism adequate for OCT sampling rate (100 kHz)
Notes: OCT systems endoscopy: fiber based
Same components but now probe goes into body
Practical requirements: Low NA 0.02

## Slide 7: Fiber scanner Fourier plane scanner
Fourier plane
Notes: Conventional microscopy: translation – angle - translation
Fourier plane scanner
-> How to achieve this angle scanning?

## Slide 8: Fiber scanner Piezoelectric actuation + Mechanical amplification
Notes: Piezoelectric actuation + Mechanical amplification
Piezoelectric tube: 4 electrodes, fixed at base
Excites a cantilever formed by fiber+GRIN
Sinusoidal excitation at fres -> 1000x mechanical gain
Resonance: mechanics define fres
-> Lower fres as much as possible: analytical calculation

## Slide 9: Fiber scanner Resonance frequency estimation
COMSOL Simulation:
762 Hz
Analytical Calculation
770 Hz
Notes: Estimation by modelling the GRIN as point weight
Dependance: Dfiber, length, mass
  Reduced d fiber to 80 um
  Placing GRIN at tip
-> Optical analysis of the scanner

## Slide 10: Optical performance Geometrical optics
Parameter | Value
NAfiber | 0.18
fGRIN | 0.91 mm
dbeam | 0.3 mm
fobj | 7.5 mm
NAOCT | 0.022
Notes: Geometrical optics:
Interrelated parameters
Set NA
Calculate resolution of such system
  OCT: Illuminate with laser, collect light

## Slide 11: Optical performance Fourier optics: illumination PSF
Infinitesimal
fiber core
Gaussian mode fiber core
PSFoptics(x)
PSFoptics(x) * MFD(x)
GRIN
objective
GRIN
objective
Notes: Start considering illumination: projector
  Fiber - lens - aperture – sample
  If fiber is point source – PSF shape defined by optics (NA), beam diameter
  If fiber has an extended mode shape: convolution of both
-> What happens with backscattered light?

## Slide 12: Optical performance Fourier optics: illumination + collection
illumination
detection
23.3 lp/mm
43 µm
CRayleigh
sample
probability
of
illumination
probability
of
detection
×
Notes: Helmholz reciprocity: Folded optical system
In order that photon reaches the detector:
  Illuminate inside PSFill
  Be accepted into the optical system
  Same PSF squaring
Squaring of PSF – Autocorrelation of MTF
-> What happens for beams at an angle? Aberrations?

## Slide 13: Optical performance ZEMAX simulation: Geometrical raytracing
spot through focus
raytracing
1.5 mm
= Airy disk (86 µm diameter)
Notes: 1.5 mm FOV
Central/marginal rays
Converge in focal plane?
Spherical lens: Spherical aberration: Marginal rays spread
-> Analyzed, simulated -> Test its behavior

## Slide 14: Implementation Single modality probe
Notes: Implementation of single modality, OCT probe
Equivalent optics, mechanics
3D printed housing
Custom PI interconnects

## Slide 15: Implementation Single modality probe: size
3D Printing by:
Laboratory for Materials Processing (Bilal Khatri)
Laboratory for the Design of Microsystems
Notes: Implementation of single modality, OCT probe
Equivalent optics, mechanics
3D printed housing
Custom PI interconnects

## Slide 16: Experimental evaluation Laser scanned imaging
1.2 mm
Notes: Experimental evaluation: Obtain Laser Scanned Images
Measure intensity of backreflected light
Setup
2D area: spiral scanning
FOV

## Slide 17: Experimental evaluation Whirling

## Slide 18: Experimental evaluation Distortion correction
Whirl patterns
Distortion Correction
Notes: Distortion:
Non ideal beam: 2 axes of symmetry: eigendirection
Cross coupling of energy between them
Whirling
Solution:
  WP far from resonance but enough gain
  Look up table measured with PSD

## Slide 19: Experimental evaluation Resolution and field of view
Lateral resolution: 48 µm
Depth of Field: 3.5 mm
Edge Spread Function
MTF
Translated mirror
,
Notes: Evaluate resolution:
Lateral resolution:
  Edge (derivative) PSF (fourier) MTF

## Slide 20: Experimental evaluation OCT cylindrical cross section imaging
Human fingertip
Human colon polyp
1.4 mm
SS-OCT System from Leitgeb Lab
1340 nm
37 nm
27 µm
Notes: Collaboration Leitgeb Lab
Acquired cylindrical cross sections
  Fingertip: Surface topology + layers
  Polyp: structures, glass
Low SNR: OCT setup not optimized for this scanning speed
Future: Assemble cylinders to 3D

## Slide 21: Conclusion and Outlook
OCT scanning engine
Amongst most compact in literature
Reduction of imaging distortions
OCT imaging
Internal structures
High penetration depth
Outlook:
Integration in double modality scanner
Combination of other scanning + full-field modalities
Accurate optical analysis
Guideline for future fiber scanners
Notes: Acknowledge Bilal Khatri (Hannemann group) and Woias group for the help in the 3D printing

## Slide 22: Compact OCT scanner for a bimodal endoscopic probe
Sergio Vilches
Notes: [Keep whole laser beam]

## Slide 23: Piezoelectric tube actuator

## Slide 24: Implementation Electrical contacts

## Slide 25: Confocal resolution

## Slide 26: Non-telecentric scanning

## Slide 27: ZEMAX raytracing of bimodal probe

## Slide 28: Sources of optical backreflections inside the probe

## Slide 29: Fiber-GRIN Bonding

## Slide 30: Spiral scanning Sampling

## Slide 31: Implementation Performance

## Slide 32: Effect of cantilever length in GRIN angle